ARIA. Side Stream Setup

- Insert dummy collection tubes into the required collection block.

- Install the collection block onto the instrument.

NOTE: If you are using a water-cooled block, please turn off the waterbath momentarily to prevent water squirting when disengaging the fittings. Turn the waterbath back on after the installation is complete.

- Turn on the deflection plates with the Side Stream window. The red warning light indicates that the plates are charged. DO NOT TOUCH THE PLATES WHILE THE RED LIGHT IS ON. THE PLATES DO NOT TURN OFF EVEN IF THE SORTING BLOCK DOOR IS OPEN.

- In  DIVA, Tturn on Test Sort and adjust the diode laser to brighten the streams if you are not able to see them.

- Open the door to the sorting chamber by loosening the thumbscrew.

- Retract the waste drawer with the waste drawer button on the Side Stream window in DIVA. This enables you to check the deflection of side streams into your tubes.

- Aim the side streams into the bottom of each collection tube with the sliders corresponding to each stream.

- Bring the waste drawer back out in DIVA and then securely close the door to the sorting chamber.

- Turn off test sort and the deflection plates.